## Sheet1
AUTOCAL SOLUTIONS PVT. LTD.
Laboratory: 06, Gurusampada Society, Vanvihar Colony,
Behind Bhonsala School, Nashik - 422 005, Maharashtra, India.
Email Id: nashiklab@autocal.net Web Adderess: www.autocal.net
Phone No.: (0253) 2352168
Cert. No.: C-1505
|  |  |  | CALIBRATION CERTIFICATE | [7 empty] | 5.10/R/E-01,Rev. 01
Page 1 of 2
Certificate No. |  |  | 16174766/N/ET/01 |  |  |  | Issue Date |  |  |  | 17/02/2017
Date of Calibration |  |  | 16/02/2017 |  |  |  | Recommended Next Calibration Date |  |  |  | 15/02/2018
Service Request Form No. |  |  |  |  | 16174766
Calibrated for |  |  |  |  | M/s.Enam Casting Pvt Ltd
D-72,MIDC Area, Satpur,Nashik - 422 007.
Date of Receipt of Instrument |  |  |  |  | 16/02/2017
Condition of Instrument on Receipt |  |  |  |  | Working |  |  | Calibrated at |  |  | Site
Details of Unit Under Calibration (UUC) :
Description |  |  | Analog H.V. Tester (AC) |  |  |  | Range |  |  | 0 to 100 kV AC @ 50Hz
Make / Model |  |  | Ericon Transformer |  |  |  | Resolution |  |  | 2 kV AC
Serial No. |  |  | HVT-726 |  |  |  | Accuracy |  |  | ±5% Fsd
ID No. |  |  | --- |  |  |  | Work Instruction No. |  |  | WI/ASPL/CAL/ET-06 &19
Environmental Conditions :
Temperature |  |  |  | 25 ±5°C |  |  |  | Relative Humidity |  |  | 60 ±15%RH
Details of Standard Used for Calibration (Traceable to National / International Standards)
Description |  |  | Make / Model |  |  | Serial No. |  |  | Traceability ( Certificate No. ) |  |  | Valid Upto
H.V. Divider with kV meter |  |  | Wavetek / 9100 |  |  | STD/NSK/ET-01 |  |  | ASPL, Vasai |  |  | 23/02/2017
( 15164490/ET/11 )
Measurement Result Refer to Page No. 2
Remarks
1. The reported expanded uncertainty of measurement is stated as a standard uncertainty of measurement multiplied by coverage factor 'k' such that the coverage probabilty corresponds to approximately 95.45% confidence level.
2. The measured values mentioned are the average of 5 readings.
3. Measured readings have been mentioned.
4. Readings taken as per customer request.
Notes
1. This report refers only to particular item calibrated at site/ Laboratory.
2. The calibration results reported are valid at the time of and under the stated conditions of measurement.
3. Any correction in this certificate invalidates this certificate.
4. This certificate shall not be reproduced except, in full without the written approval from the Laboratory.
Calibrated By | [9 empty] | Authorised Signatory
Bhupendra Ozarkar | [9 empty] | Baviskar P.U.
Calibration Engineer | [9 empty] | Technical Manager
AUTOCAL SOLUTIONS PVT. LTD.
Laboratory: 06, Gurusampada Society, Vanvihar Colony,
Behind Bhonsala School, Nashik - 422 005, Maharashtra, India.
Email Id: nashiklab@autocal.net Web Address: www.autocal.net
Phone No.: (0253) 2352168
5.10/R/E-01,Rev. 01
Page 2 of 2
Certificate No. |  |  | 16174766/N/ET/01 |  |  |  | Issue Date |  |  |  | 17/02/2017
Date of Calibration |  |  | 16/02/2017 |  |  |  | Recommended Next Calibration Date |  |  |  | 15/02/2018
Measurement Result | [13 empty] | UNCERTAINTY EVALUATION
Range | Test Point |  | Nominal Value on Standard |  | Measured Value on UUC |  | Deviation |  | Deviation (%Fsd) | (±) Expanded Uncertainty (%Fsd) |  | Coverage Factor (k) |  | Analog H.V. Tester (AC)
[14 empty] | Certificate No. |  | 0
Parameter : High Voltage AC @ 50 Hz | [12 empty] | Test Point | Type "A" (Repeatability) |  |  |  |  | Avg. | SD | Ua | Limits (%) | Divisor | Ub1 | Limits (%) | Divisor | Ub2 | Limits (A) | Divisor | Ub3 (%) | Uc (%) | DOF | k | Ux (%) | CMC as per Scope (%rdg)
0 to 100kV AC @ 50Hz | 0 | kV | 0 | kV | 0 | kV | 0 | kV | --- | --- |  | --- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 0.026 | 2 | 0.013 | 2 | 1.732 | 1.1547 | 2 | 1.732 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0.17 | #DIV/0!
| 10 |  | 10.8 |  | 10 |  | -0.8 |  | -0.8 | 11.7761 |  | 2 | 10 | 10.8 | 10.8 | 10.8 | 10.81 | 10.8 | 10.802 | 0.0044721 | 0.02 | 0.026 | 2 | 0.013 | 2 | 1.732 | 1.1547 | 2 | 1.732 | 5.7737 | 5.8881 | 30048956302.13 | 2 | 11.7761 | 0.17 | 11.7761
| 20 |  | 21.89 |  | 20 |  | -1.89 |  | -1.89 | 6.2186 |  | 2 | 20 | 21.89 | 21.89 | 21.89 | 21.87 | 21.89 | 21.886 | 0.0089443 | 0.02 | 0.026 | 2 | 0.013 | 2 | 1.732 | 1.1547 | 2 | 1.732 | 2.8868 | 3.1093 | 2336660279.834 | 2 | 6.2186 | 0.17 | 6.2186
| 30 |  | 32.94 |  | 30 |  | -2.94 |  | -2.94 | 4.4889 |  | 2 | 30 | 32.94 | 32.94 | 32.94 | 32.95 | 32.94 | 32.942 | 0.0044721 | 0.0066667 | 0.026 | 2 | 0.013 | 2 | 1.732 | 1.1547 | 2 | 1.732 | 1.9246 | 2.2444 | 51388161248.5207 | 2 | 4.4889 | 0.17 | 4.4889
| 40 |  | 43.72 |  | 40 |  | -3.72 |  | -3.72 | 3.6971 |  | 2 | 40 | 43.72 | 43.72 | 43.72 | 43.73 | 43.72 | 43.722 | 0.0044721 | 0.005 | 0.026 | 2 | 0.013 | 2 | 1.732 | 1.1547 | 2 | 1.732 | 1.4434 | 1.8485 | 74728362991.6618 | 2 | 3.6971 | 0.17 | 3.6971
| 50 |  | 54.17 |  | 50 |  | -4.17 |  | -4.17 | 3.2662 |  | 2 | 50 | 54.17 | 54.17 | 54.17 | 54.18 | 54.17 | 54.172 | 0.0044721 | 0.004 | 0.026 | 2 | 0.013 | 2 | 1.732 | 1.1547 | 2 | 1.732 | 1.1547 | 1.6331 | 111139567401.5154 | 2 | 3.2662 | 0.17 | 3.2662
| 60 |  | 64.42 |  | 60 |  | -4.42 |  | -4.42 | 3.0064 |  | 2 | 60 | 64.42 | 64.42 | 64.42 | 64.43 | 64.42 | 64.422 | 0.0044721 | 0.0033333 | 0.026 | 2 | 0.013 | 2 | 1.732 | 1.1547 | 2 | 1.732 | 0.96228 | 1.5032 | 165423554675.8315 | 2 | 3.0064 | 0.17 | 3.0064
| 70 |  | 74.61 |  | 70 |  | -4.61 |  | -4.61 | 2.8382 |  | 2 | 70 | 74.61 | 74.61 | 74.61 | 74.62 | 74.61 | 74.612 | 0.0044721 | 0.0028571 | 0.026 | 2 | 0.013 | 2 | 1.732 | 1.1547 | 2 | 1.732 | 0.82481 | 1.4191 | 243449167279.6412 | 2 | 2.8382 | 0.17 | 2.8382
Calibrated By | [9 empty] | Authorised Signatory
Bhupendra Ozarkar | [9 empty] | Baviskar P.U.
Calibration Engineer | [9 empty] | Technical Manager
********** End of Certificate **********
Note on X77: user: Stadard's Calibrator Uncertainty taken from Cal. Certificate
Note on AA77: user: Standard's Calibrator Accuracy taken from Manf. Specification
Note on AE77: user: UUC Resolution

## Sheet2
| Certificate No. |  |  | 16174766/N/ET/01 |  |  |  | Issue Date |  |  |  |  | 17/02/2017
| Date of Calibration |  |  | 16/02/2017 |  |  |  | Recommended Next Calibration Date |  |  |  |  | 15/02/2018
| Service Request Form No. |  |  |  |  | 16174766
| Calibrated for |  |  |  |  | M/s.Enam Casting Pvt Ltd
D-72,MIDC Area, Satpur,Nashik - 422 007.
UNCERTAINTY EVALUATION
Analog H.V. Tester (AC)
| Certificate No. |  | 16174766/N/ET/01
Test Point | Type "A" (Repeatability) |  |  |  |  | Avg. | SD | Ua | Limits (%) | Divisor | Ub1 | Limits (%) | Divisor | Ub2 | Limits (A) | Divisor | Ub3 (%) | Uc (%) | DOF | k | Ux (%) | CMC as per Scope (%rdg) | Final Un.
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 0.026 | 2 | 0.013 | 2 | 1.732 | 1.1547 | 2 | 1.732 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0.17 | #DIV/0!
10 | 10.8 | 10.8 | 10.8 | 10.81 | 10.8 | 10.802 | 0.0044721 | 0.02 | 0.026 | 2 | 0.013 | 2 | 1.732 | 1.1547 | 2 | 1.732 | 0.57737 | 1.2913 | 69500020.8523 | 2 | 2.5825 | 3 | 3
20 | 21.89 | 21.89 | 21.89 | 21.87 | 21.89 | 21.886 | 0.0089443 | 0.02 | 0.026 | 2 | 0.013 | 2 | 1.732 | 1.1547 | 2 | 1.732 | 0.57737 | 1.2913 | 69500020.8523 | 2 | 2.5825 | 3 | 3
30 | 32.94 | 32.94 | 32.94 | 32.95 | 32.94 | 32.942 | 0.0044721 | 0.0066667 | 0.026 | 2 | 0.013 | 2 | 1.732 | 1.1547 | 2 | 1.732 | 0.57737 | 1.2911 | 5627100984.858 | 2 | 2.5822 | 3 | 3
40 | 43.72 | 43.72 | 43.72 | 43.73 | 43.72 | 43.722 | 0.0044721 | 0.005 | 0.026 | 2 | 0.013 | 2 | 1.732 | 1.1547 | 2 | 1.732 | 0.57737 | 1.2911 | 17784003037.6468 | 2 | 2.5822 | 3 | 3
50 | 54.17 | 54.17 | 54.17 | 54.18 | 54.17 | 54.172 | 0.0044721 | 0.004 | 0.026 | 2 | 0.013 | 2 | 1.732 | 1.1547 | 2 | 1.732 | 0.57737 | 1.2911 | 43417507335.3314 | 2 | 2.5822 | 3 | 3
60 | 64.42 | 64.42 | 64.42 | 64.43 | 64.42 | 64.422 | 0.0044721 | 0.0033333 | 0.026 | 2 | 0.013 | 2 | 1.732 | 1.1547 | 2 | 1.732 | 0.57737 | 1.2911 | 90030015121.4758 | 2 | 2.5822 | 3 | 3
70 | 74.61 | 74.61 | 74.61 | 74.62 | 74.61 | 74.612 | 0.0044721 | 0.0028571 | 0.026 | 2 | 0.013 | 2 | 1.732 | 1.1547 | 2 | 1.732 | 0.57737 | 1.2911 | 166791127916.5786 | 2 | 2.5822 | 3 | 3
Note on J16: user: Stadard's Calibrator Uncertainty taken from Cal. Certificate
Note on M16: user: Standard's Calibrator Accuracy taken from Manf. Specification
Note on Q16: user: UUC Resolution